## Slide 1: The measurement of object distance for projected virtual objects using a scanning pentaprism and digital alignment telescope
Kevin Sweeney SPIE Photonics West
David A. Imrie, PhD OPTO 13373-8

## Slide 2: An imager test scenario must be with the use case.
Conjugate Distance
Aperture Size
Wavelength Spectrum
consistent
Target Shape
How do we ensure
regardless of ,
or ?
OPTO 13373-8

## Slide 3: The Scanning Pentaprism has the versatility we need.
 | Restriction |  |  |  | 
Method | Pinhole Source | Monochromatic Source | Instrument EFL >> Collimator EFL | Relay Optic > Collimator Aperture | Accommodating Features in Collimator
Laser Unequal Path Interferometer | ✖ | ✖ | ✓ | ✖ | ✖
Calibrated Camera | ✓ | ✓ | ✖ | ✖ | ✓
Shack-Hartmann Wavefront Sensor | ✖ | ✓ | ✓ | ✖ | ✓
Shear Plate Interferometer | ✖ | ✖ | ✓ | ✖ | ✓
Autocollimation | ✓ | ✓ | ✓ | ✖ | ✖
Scanning Pentaprism | ✓ | ✓ | ✓ | ✓ | ✓
There are several methods for ensuring collimation…
OPTO 13373-8

## Slide 4: A pentaprism provides a constant deflection angle.
90°
90°
Rotate Prism
Deflection angle is unchanged!
Incoming light ray
Deflected light ray
Incoming light ray
Deflected light ray

## Slide 5: The Scanning Pentaprism Approach
Alignment telescope and collimator are rigidly affixed to each other.
  Moving only the pentaprism drastically reduces straightness required of the linear guideway
Location of the telescope image is proportional to angle of ray collected from target projector
Target Projector
Alignment Telescope + Scanning Pentaprism
Defocus
Pentaprism Position
Telescope Image Location
OPTO 13373-8
Produces
Ray Fan Plot

## Slide 6: Geometrical Optics Analysis — See Manuscript
Pentaprism = Aperture Stop → Defines line of sight in object space
Image Sensor
Real or Projected Object
Line of Sight
Line of Sight
Line of Sight
Telescope Objective
OPTO 13373-8
Image Sensor
Alignment Telescope Objective Lens
Scanning Pentaprism Aperture
Real or Projected Object
Image
Unfolded
Chief ray height at telescope objective
Image location
Chief ray angle in image space
Chief ray height on image sensor
  — Nominal aperture spacing
  — Telescope objective focal length
  — Pentaprism position
  — Chief ray height on sensor
  — Object distance
Known by design
Monitored during operation
Quantity to be calculated

## Slide 7: The Scanning Pentaprism Demonstrator
Scan Range
350 mm
Weight
16.9 kg
Pixel Subtense
2 arcsec
Full Field of View
1.38°
Spectral Range
300-1000nm
OPTO 13373-8

## Slide 8: Compensating Spherical Aberration
OPTO 13373-8
Focus Setting Adjustment
Autocollimator Crosshair in Live Video
Measured Ray Fans

## Slide 9: Measuring Departure from Collimation
Laser unequal path interferometer (LUPI) is both:
  Point source for collimation measurement
  Independent measurement of collimation
OPTO 13373-8
LUPI
350mm Reflective Collimator
Return Flat for LUPI
Scanning Pentaprism
Projected Distance > 35km
Projected Distance = 34km
Projected Distance = 23km

## Slide 10: Measuring Finite Conjugates
OPTO 13373-8
The scanning pentaprism is a versatile technique for determining conjugate distance of any optical beam!
White Light Point Source Generator
Surveyor’s Tape Measure

## Slide 11: Measuring Parallax
OPTO 13373-8
The reticles of telescopic sights can be measured provided uniform background illumination.
For the case of rifle scopes, this can directly measure the “zero parallax distance” on the objective side of the scope.

## Slide 12: Thank you!
Lucas Delbanco
Design and build of the demonstrator
Michael Halligan
Design and build of demonstrator tip/tilt base
Dr. Jian Zhang
Software for motion control, image acquisition, and centroid tracking
Dr. Stephen D. Fantone
Funding and other resources at Optikos Corporation
Wetthauer and Brodhun, "Objektiv Untersuchungen," Zeitschrift für Instrumentenkunde, vol. 40, no. 5, pp. 96-97 (May 1920)
R. N. Wilson, ""Matching Error" (Spherical Aberration) in the Hubble Space Telescope (HST): Some Technical Comments," The Messenger, no. 61, pp. 22-24 (1990)
W. Scott Smith "Testing the 8.3-meter telescope optics", Proc. SPIE 2576, International Conference on Optical Fabrication and Testing, pp. 220-231 (August 1995); https://doi.org/10.1117/12.215597
P. Su, J. H. Burge, B. Cuerden, J. Sasian and H. M. Martin, "Scanning pentaprism measurements of off-axis aspherics," Proc. SPIE 7018, Advanced Optical and Mechanical Technologies in Telescopes and Instrumentation, 70183T (25 July 2008); https://doi.org/10.1117/12.789588
S. Qian, W. Jark, P. Z. Takacs, K. J. Randall and W. Yun, "In situ surface profiler for high heat load mirror measurement," Optical Engineering 34(2), pp. 396-402 (1 February 1995); https://doi.org/10.1117/12.194834
S. Qian, W. Jark and P. Z. Takacs, "The penta‐prism LTP: A long‐trace‐profiler with stationary optical head and moving penta prism," Review of Scientific Instruments, vol. 66, p. 2562–2569 (1995); https://doi.org/10.1063/1.1145658
K. von Bieren, "Pencil beam interferometer for aspherical optical surfaces," Proc. SPIE 0343 Laser Diagnostics, pp. 101-108 (19 November 1982); https://doi.org/10.1117/12.933743
P. C. V. Mallik, C. Zhao and J. H. Burge, "Measurement of a 2-meter flat using a pentaprism scanning system," Optical Engineering 46(2), 023602 (1 February 2007); https://doi.org/10.1117/1.2700386
Optics and optical instruments — Accuracy of optical transfer function (OTF) measurement, ISO 11421:1997(E), Geneva, Switzerland, (September 15 1997)
K. Mohammadi, Y. S. Lee and T. Jennewein, "Characterization of optical aberrations with scanning pentaprism for large collimators," CLEO: Applications and Technology, Imaging Techniques Across Multiple Modalities and Dimensions (ATh1D), San Jose, 2022.
Acknowledgments
References
OPTO 13373-8
Come see the scanning pentaprism instrument at Optikos booth #841